## tbl_4c
Table 4c
Uranium and Thorium Mass Fraction (NDA)
Uncertainty (% rel., k=1)
Method | Technique | Instrument* | Material (homogeneous) | Element | Mass fraction range | u(r) | u(s) | u(c) | Notes
HRGS | T1 | XMCG | Powder, solution, slurry | NU | 1 % to 25 % | 2 | 3 | 3.6056 | DX, IT, KE, LZ, EG=185.7, WT=5, LT=1000, FLD, HY
MRGS | T1 | XMCL | Powder, solution, slurry | NU | 1 % to 25 % | 2 | 3.5 | 4.0311 | IT, KE, LZ, EG=185.7, WT=5, LT=1000, FLD, HY
|  | XMCC | Powder, solution, slurry | NU | 1 % to 25 % | 5 | 4 | 6.4031 | IT, KE, LZ, EG=185.7, WT=5, LT=1000, FLD, HY
HRGS | T1 | XMCG | Powder, solution, slurry | DU | 1 % to 25 % | 13.6317 | 3.5451 | 14.0851 | D1, D2, IT, RE, LZ, EG=185.7, ISO-LO, WT=1, LT=1000, FLD, HY
|  | XMCG | Powder, solution, slurry | NU | 1 % to 25 % | 7.4161 | 3.4052 | 8.1606 | D1, D2, IT, RE, LZ, EG=185.7, ISO-LO, WT=1, LT=1000, FLD, HY
|  | XMCG | Powder, solution, slurry | LEU | 1 % to 25 % | 5.7447 | 3.6524 | 6.8075 | D1, D2, IT, RE, LZ, EG=185.7, ISO-LO, WT=1, LT=1000, FLD, HY
|  | XMCG | Powder, solution, slurry | DU | 1 % to 25 % | 17.5292 | 3.2303 | 17.8244 | D2, D3, IT, RE, LZ, EG=185.7, ISO-HI, WT=1, LT=1000, FLD, HY
|  | XMCG | Powder, solution, slurry | NU | 1 % to 25 % | 16.6152 | 3.0473 | 16.8923 | D2, D3, IT, RE, LZ, EG=185.7, ISO-HI, WT=1, LT=1000, FLD, HY
|  | XMCG | Powder, solution, slurry | LEU | 1 % to 25 % | 13.5828 | 3.0334 | 13.9174 | D2, D3, IT, RE, LZ, EG=185.7, ISO-HI, WT=1, LT=1000, FLD, HY
| T2 | XMCG | Powder, slurry, sludge | NU | > 25 % | 2 | 0.7 | 2.119 | DX, TS, KE, SM=2, EG=185., LT=1000, TD
|  | XMCG | Solution (organic and aqueous) | NU | 1 % to 25 % | 1.8 | 0.8 | 1.9698 | DX, TS, KE, SM=10 mL, EG=185.7, LT=1000, TD
|  | XMCG | Powder | LEU | > 25 % | 1.4 | 0.4 | 1.456 | D2, TS, RE, SM=50 to 100, EG=1001, ISO-LO, LT=2000, TD
|  | XMCG | Powder, ore | Th | 0.1 % - 1 % | 5 | 4 | 6.4031 | D2, D3, TS, SM=10, EG=583/911, LT=3000, FLD, TD
|  | XMCG | Powder, ore | Th | > 1 wt% | 1.5 | 1 | 1.8028 | D2, D3, TS, SM=10, EG=583/911, LT=1000, FLD, TD
| T3 | XMCG | Powder, ore | Th | 0.1 % - 1 % | 0.8 | 4.4 | 4.4721 | D2, D3, TS, RE, EG=583/911/2614, PA-911=30,000, LAB, TD
|  | XMCG | Powder, ore | Th | > 1 % | 0.6 | 1.5 | 1.6155 | D2, D3, TS, RE, EG=583/911/2614, PA-911=30,000, LAB, TD
|  | XMCG | Powder, solution | LEU | > 25 % | 13 | 3 | 13.3417 | D2, D3, TS, RE, SM>=150, EG=185.7/1001, LT=1000, FLD, TD
XRF | T4 | ED-XRF | Powder, solution | U/Th | 0.1 % - 25 % | 3 | 3 | 4.2426 | XRF Literature listed in Narrative
|  | WD-XRF | Solution (organic and aqueous) | U | 1 % - 25 % | 3 | 3 | 4.2426 | Refinery Plant Report
| T5 | WD-XRF | Powder, ore | Th | 0.1 % - 10 % | 3.3 | 6 | 6.8476 | SGAS-NML Measurements
*click a hyperlink to open method's narrative
Uncertainty categories:
Prospective Values: are presented in blue font
Assumptions:
1 | Only homogeneous material (no sampling error included)
2 | For method T1, systematic uncertainty includes sample matrix, calibration, wall thickness determiantion and detector specific spectrum fitting components
Method categories:
HRGS | High Resolution Gamma Spectrometry
MRGS | Medium Resolution Gamma Spectrometry
XRF | X-Ray Fluorescent Analysis
Technique categories:
T1 | Concentration meter principle (in-situ measurements)
T2 | Absolute gamma-spectrometry calibrated against reference materials (on-site laboratory setup)
T3 | In-situ absolute gamma-spectrometry calibrated mathematically (e.g. by Monte Carlo or ISOCS) in situations when no appropriate calibration standard is available
T4 | XRF using a calibration curve
T5 | XRF using fundamental parameters
Notes:
EG | Assay gamma ray energy (keV)
IT | Infinitely thick samples (≥ 5 mfp at the assay gamma ray energy)
TS | Thin samples (<< 1 mfp at the assay gamma ray energy)
KE | Material with known enrichment (e.g., NU at concentration, refinery and conversion plants)
LZ | Low-Z matrix, known contents of high-Z additives
RE | Radioactive equilibrium. In the case of uranium, enrichment can be measured using a complementary peak ratio method (MGAU, FRAM or similar)
SM | Sample mass (g)
WT | Container wall thickness (mm)
LT | Measurement live time (s)
ISO-LO | Low-energy isotopic analysis (60 keV to 250 keV)
ISO-HI | High-energy isotopic analysis (120 keV to 1010 keV)
PA-XXX | Number of counts (area) of the peak at energy XXX keV
LAB | Laboratory measurements
FLD | Field measurements
Instrument categories:
XMCG | HPGe detectors, including both electrically or LN2 cooled, coupled to any MCA
XMCC | CZT, active volume ≥ 1500 mm3, collimator: Æ15 x 5 mm, any MCA
XMCL | LABR, crystal Æ50 x 13 mm (or larger), collimator: Æ44 x 20 mm, any MCA
XMCG types:
D1 | Energy resolution < 750 eV @ 122 keV. Active area > 500mm2. Thickness > 10mm. Typical detectors in XMCG-D1 category: planar and semi-planar detectors (GL1015, BE1015, BE2830, BE3830, BE5030, Profile S/SP, GLP).
D2 | Energy resolution < 1.0 keV @ 122 keV and < 2.0 keV @ 1332 keV. Relative efficiency > 15%. Typical detectors in XMCG-D2 category: semi-planar and small volume coaxial detectors (BE2830, BE3830, BE5030, Profile S/SP/C/M, GMX, Falcon, Aegis).
D3 | Energy resolution < 1.5 keV @ 122 keV and < 2.3 keV @ 1332 keV. Relative efficiency > 15%. Typical detectors in XMCG-D3 category: portable electically cooled coaxial detectors and large volume coaxial detectors (Micro-Detective (ECGS), Aegis, GC9023, Profile C/M, GMX).
Method used to derive ITVs:
TD | Top-down
BU | Bottom-up
HY | Hybrid