## tbl_6b
Table 6b
Pu Isotopic Abundances (NDA)
[6 empty] | 239Pu |  | 240Pu |  | 240Pueff |  | Specific power
Method | Technique | Instrument* | Material | Burnup | Shielding | (% rel., k=1) |  | (% rel., k=1) |  | (% rel., k=1) |  | (% rel., k=1) |  | Notes
[6 empty] | u(r) | u(s) | u(r) | u(s) | u(r) | u(s) | u(r) | u(s)
HRGS | T1, LE | XMCG-D1 | Pu | L | L | 0.14406 | 0.066005 | 1.5395 | 0.79238 | 1.5318 | 0.7873 | 1 | 0.24106 | 1
|  |  | Pu | H | L | 0.6 | 0.3 | 1.2 | 0.4 | 0.86938 | 0.3 | 0.56837 | 0.26488 | 1, 3
| T2, LE | XMCG-D1 | Pu | H | L | 0.7 | 0.3 | 1.2 | 0.6 | 1.5 | 0.5 | 0.6 | 0.3 | 1, 3
| T1, LE | XMCG-D1 | MOX | H | L | 1 | 0.4 | 2.6 | 0.8 | 2.6 | 0.5 | 1 | 0.7 | 6
| T1, HE | XMCG-D2 | Pu | L | L/H | 0.61015 | 0.17569 | 6.2142 | 1.6665 | 6.1708 | 1.6529 | 1.0469 | 0.42997 | 2
|  |  | Pu | H | L/H | 2.059 | 0.61823 | 6.2329 | 1.735 | 5.567 | 1.5501 | 1.5426 | 0.94985 | 2
| T1, HE | XMCG-D3 | Pu | L | L/H | 1.6698 | 0.2237 | 16.0858 | 3.0563 | 15.7217 | 3.0324 | 2.3948 | 0.38975 | 7, 11
|  |  | Pu | H | L/H | 4.9026 | 4.0823 | 15.1881 | 9.8954 | 11.2626 | 6.7118 | 3.3107 | 0.6567 | 7, 11
|  | XMCG-D3 | MOX | H | L | 2.6 | 1.9 | 6.7 | 6 | 3.4 | 2.7 | 2.7 | 3.1 | 10
MRGS | T3/T4, HE | XMCC | Pu | L | L/H | 2.5 | 1.5 | 20 | 15 | N/A | N/A | N/A | N/A | 4, 8
|  |  | Pu | H | L/H | 3.5 | 2.5 | 15 | 10 | N/A | N/A | N/A | N/A | 4, 8
| T3/T4, HE | XMCL | Pu | L | L/H | 2.5 | 2 | 20 | 15 | N/A | N/A | N/A | N/A | 5, 9
|  |  | Pu | H | L/H | 3.5 | 3 | 15 | 15 | N/A | N/A | N/A | N/A | 5, 9
*click a hyperlink to open method's narrative
N/A: not applicable
Uncertainty categories:
ITV: are presented in black font
Prospective Values: are presented in blue font
Assumptions / considerations:
1 | LANL data, FRAM v.6.1 analysis with parameter set GePlnr_Pu_060-230
2 | LANL data, FRAM v.6.1 analysis with parameter set GeCoax_Pu_180-1010
3 | Additional source of data from IAEA NML obtained using FRAM v.5.2 / MGA (10 analyses) 129 keV = 50000 cnts, OCR=22 kcps, LT=600 s. The prospective values were established based on both, LANL and IAEA dataset.
4 | PAT analysis, MRGS Intercomparison exercise - XMCC: 413 keV = 2000 cnts (HBPu) and 20000 cnts (LBPu), OCR=10 kcps, LT=600 s.
5 | PAT analysis, MRGS Intercomparison exercise - XMCL: 413 keV = 10000 cnts (HBPu) and 50000 cnts (LBPu), OCR=10 kcps, LT=600 s.
6 | IAEA field data, FRAM v.6.1 analysis with parameter set GePlnr_Pu_060-230, 129 keV = 30000 cnts, achievable with standard IAEA XMCG-D1 setup at OCR=10 kcps for LT=600 s
7 | LANL data obtained using two 1st generation Ortec Detective instruments. Data measured in 2005 and 2016. 28 data sets of 14 sources. 539 spectra. FRAM analysis with parameter set detective_pu_120-460_220210. Normalized to 600 s LT.
8 | LANL data, FRAM v.6.1 analysis with parameter set CZT500_Pu_120-500 with correlated 240Pu option turned on in addition to correlated 238Pu and correlated 242Pu. 9 data sets, 299 spectra. Normalized to 600 s LT.
9 | LANL datam FRAM v.6.1 analysis with parameter set LaBr_Pu_200-750 with correlated 240Pu option turned on in addition to correlated 238Pu and correlated 242Pu. 21 data sets, 530 spectra. Normalized to 600 s LT.
10 | IAEA field data, FRAM v.6.1 analysis with a modified parameter set "det_coax_120_800_2_ecgs" (interferences with Cs-137 and Ag-110 m), very high BUPu (239Pu isotope mass fraction < 55 %, 2-3 y after chemical purification, DU/Pu = 10), 129 keV = 5000 cnts, 208 keV = 700000 cnts, achievable with standard IAEA XMCG-D3 setup at OCR=5 kcps for LT=1000 s (DT < 20%)
11 | EC-JRC data, steel filter up to 16 mm, lead filter up to 4 mm, 1400 spectra in 40 sets, analyzed with FRAM v.5.2 and parameter set Pu_Cx_180-1010.
Method categories:
HRGS | High Resolution Gamma Spectrometry
MRGS | Medium Resolution Gamma Spectrometry
Technique categories:
T1 | Peak ratio, intrinsic efficiency calibration, "known" 242Pu
T2 | Peak ratio, intrinsic efficiency calibration, 242Pu by isotopic correlations
T3 | Peak ratio, intrinsic efficiency calibration, 238Pu and 242Pu by isotopic correlations
T4 | Peak ratio, intrinsic efficiency calibration, 238Pu, 240Pu, and 242Pu by isotopic correlations
LE | Analysis in the low energy range < 230 keV
HE | Analysis in the high energy range > 120 keV
Instrument categories:
XMCG | HPGe detector coupled to any MCA, including portable equipment. Electrically or LN2 cooled.
XMCC | CZT, >=500 mm3, coupled to any MCA
XMCL | LABR, thickness ≥ 12.7 mm, diameter ≥ 25.4 mm, coupled to any MCA
Material categories:
Pu | PuO2 clean powder, long time after chemical separation > 1 year
MOX | Mixed oxide clean powder based on reactor grade Pu (RGPu), pellets, rods, assemblies: 10 < NU-DU/RGPu < 100
Geometries (Shielding) categories:
L | Low | < 3 mm steel (or equivalent, intense X-ray peaks)
H | High | > 10 mm steel (or equivalent, no/low intensity X-ray peaks) and < 5 mm lead or equivalent
Burnup categories:
L | Low | 239Pu isotope mass fraction > 80 %
H | High | 239Pu isotope mass fraction ≤ 80 %
XMCG types:
D1 | Energy resolution < 750 eV @ 122 keV. Active area > 500 mm2. Thickness > 10 mm. Typical detectors in XMCG-D1 category: planar and semi-planar detectors (GL1015, BE1015, BE2830, BE3830, BE5030, Profile S/SP, GLP).
D2 | Energy resolution < 1.0 keV @ 122 keV and < 2.0 keV @ 1332 keV. Relative efficiency > 15 %. Typical detectors in XMCG-D2 category: semi-planar and small volume coaxial detectors (BE2830, BE3830, BE5030, Profile S/SP/C/M, GMX, Falcon, Aegis).
D3 | Energy resolution < 1.5 keV @ 122 keV and < 2.3 keV @ 1332 keV. Relative efficiency > 15 %. Typical detectors in XMCG-D3 category: portable electically cooled coaxial detectors and large volume coaxial detectors (Micro-Detective (ECGS), Aegis, GC9023, Profile C/M, GMX).